## Uputstvo
U P U T S T V O : Ponuđač popunjava Prilog B konkursne dokumentacije za javnu nabavku sitna laboratorijska oprema, po partijama, unošenjem traženih podataka u odgovarajuća polja/kolone u narednom listu (sheet-u) ovog fajla (- Obrazac ponude sa strukturom cene - obrazac 1 tačka 5) - opis predmeta nabavke - Materijal za obrazovanje i nauku -- sitna laboratorijska oprema, po partijama). Ponuđač, u okviru jedne stavke, može da ponudi jedno ili više dobara. Za svako od ponuđenih dobara ponuđač je dužan da unese tražene podatke (naziv ponuđenog dobra, šifru i naziv proizvođača) Način unosa cene: Ponuđač unosi samo jedničnu cenu bez PDV po jedinici mere (kolona:VII ). Ponuđač unosi samo jednu jediničnu cenu bez PDV po jedinici mere, zaokruženu na dve decimale, bez obzira na broj ponuđenih dobara (naziva ponuđenih dobara) u okviru jedne stavke. Nije potrebno unositi vrednosti iz ostalih kolona (Ukupna cena bez PDV/Iznos PDV (nominalno)/Ukupna cena sa PDV, kao ni ukupnu vrednost ponude sa i bez PDv i iznos PDV), koje se same obračunavaju prema unapred zadatim formulama. Kao stopa PDV-a, koje je uračunata/zadata u formuli, je stopa od 20%. Ako se konstatuje računska greška, ista će biti otklonjena rukovodeći se jediničnom cenom. Ponuđač obrazac mora da popuni, overi pečatom i potpiše, čime potvrđuje da su tačni podaci koji su u navedeni. Ukoliko ponuđači podnose zajedničku ponudu, predmetni obrazac se potpisuje I overava u skladu sa sporazumom. Ponuđač je dužan da: - dostavi Prilog B konkursne dokumentacije za javnu nabavku sitne laboratorijske opreme, po partijama - Obrazac ponude sa strukturom cene - obrazac 1 tačka 5) - opis predmeta nabavke popunjen, odštampan i potpisan; - dostavi predmetni Prilog i u elektronskom obliku (excel fajl), na CD/DVD-u ili USB, nepotpisanu kopiju. U slučaju neslaganja između podataka (uključujući i cene) u štampanom obliku i kopije dostavljene u elektronskom obliku, verodostojnom će se smatrati štampana verzija. Ponuđač je dužan da unese i podatke koji se odnose na rok isporuke, rok važenja ponude i garantni rok.

## sitna lab. oprema
PRILOG B KONKURSNE DOKUMENTACIJE ZA JAVNU NABAVKU - OBRAZAC PONUDE SA STRUKTUROM CENE - OBRAZAC 1 TAČKA 5) OPIS PREDMETA NABAVKE Materijal za obrazovanje i nauku - SITNA LABORATORIJSKA OPREMA, po partijama
I -Stavka | II - Naziv dobra | III - Poseban zahtev, šifra | IV -Naziv ponuđenog dobra, proizvođač, šifra | V- Јedinica mere | VI -Količina | VII - Jedinična cena bez PDV po jedinici mere | VIII - Ukupna cena bez PDV | IX - Iznos PDV (nomimalno) | X- Ukupna cena sa PDV
1 | Vacuum pump | vaccum pump V-100, volt 100-240V, 50/60 Hz, power consumption 150W, power consumption in ECO mode 70 W, Terminal vacuum (absolute) 10 mbar (± 2 mbar), Pump capacity 1.5 m3/h, Brushless DC motor, Rotation speed 1280 rpm max., Noise level 32 to 57 dBA 57 dBA: with 100% load, Certificates CE, Max. altitude above sea level 2000 m, Ambient temperature 5 - 40 °C, Maximum relative humidity 80 % for temperatures up to 31 °C decreasing linearly to 50 % at 40 °C |  | kom | 1 |  | 0 | 0 | 0 | 20%
2 | SPE Manifold | SPE Vacuum Manifold, standard, 12-port model |  | kom | 1 |  | 0 | 0 | 0 | 20%
3 | pH meter | Portable meter ProfiLine pH 3110 - pH or ORP measurements, Easy 1- to 3-point calibration with adjustable calibration timer, Robust and waterproof (IP 67), Calibration memory current calibration, Display 7-Segement LCD, 3205: LCD Graphik, Power supply 4 x 1.5 V AA oder 4 x 1.2 V NiMH-rechargeable battery, Continuous operation time up to 3000 hours, pH - 2.0...20.0 ± 0.1 pH - 2.00...20.00 ± 0.01 pH - 2.000...19.999 ± 0.005 pH, mV ±1200.0 ± 0.3 mV ± (2000.0 ± 1) mV, Temperature - 5.0...105.0 °C ± 0.1 °C, Calibration 1-. 2-. 3 point WTW Technical and NIST Buffer, pH/ORP All SenTix® electrodes with DIN plug |  | kom | 1 |  | 0 | 0 | 0 | 20%
4 | pH elektroda | standard pH electrode, Connector DIN, waterproof, Cable 1 m, Int. temp. probe yes, Diameter [mm] 12, Temperature [°C] 0 - 80, UNSPSC 41115604 pH-elektroder, Qty per Salesunit 1 each |  | kom | 2 |  | 0 | 0 | 0 | 20%
5 | pH elektroda | Special pH Electrode, Temperature (Metric) Range 0 to 100°C, Temperature (Metric) Operating 0 to 100°C, Junction Platinum, Connector Types Fixed cable, Includes DIN, banana electrode plug, 1m cable, pH Range 0 to 14, Material Glass shaft, Platinum diaphragm |  | kom | 1 |  | 0 | 0 | 0 | 20%
6 | JAR tester | Flocculation Tester, 6 places, Speed can be changed during the run Illuminated back panel to simplify sample observation Stirring rods adjustable in height without tools Reliable and repeatable results from stirring quality Precise speed from an electronically controlled DC gear motor Strong resistance to chemical and mechanical corrosion Independently selectable speeds 10-15-30-45-60-90-120-150-200-300 rpm Same product can be submitted to different conditions in parallel |  | kom | 1 |  | 0 | 0 | 0 | 20%
7 | High Quality Industrial Peristaltic Pump_N6-6L Peristaltic Pump | Flow rate range 0.00166-570mL/min Speed range 1-150 rpm Speed resolution 0.01 rpm Dispensing volume range 0.01-9999.99 mL Dispensing volume resolution 0.01 mL Dispensing time 0.1-9999.99 s Pause time 0.1-9999.99 s Time resolution 0.01 s Copy numbers 1-9999 times, setting '0' means unlimited Back suction angle 0-360° Dispensing accuracy <±0.5% Motor type Steeper Motor Display Industrial grade 4.3" color LCD display Control method Touch screen and Mechanical keypad Keypad lifetime 300,000 times Start/stop, direction signal Passive switch signal, such as foot pedal switch, Active switch signal: 5V, 12V and 24V for option Power supply 220V±10% 50Hz/60Hz(Standard), 110V±10% 50Hz/60Hz(Optional) Output interface Output motor working status(Open-Collector output) Drive dimension max 261.4x157.3x236.9mm(Lx W x H) Drive weight max 4.40 kg Power consumption <50W Environment temperature 0-40℃ Relative humidity <80% IP rate IP31 |  | kom | 1 |  | 0 | 0 | 0 | 20%
8 | Sun radiation measuring device | Resolution 0,1 W/m2 , 0,1 Btu/(ft2 -h) Presicion ±10 W/m2 [±3 Btu/(ft2 -h)]or ±5%of measured value Temperature error ±0,38 W/m2 /°C [±0,12 Btu/(ft2 -h)/°C] deviation at 25°C Display 3-3/4 LCD display, the largest numerical value shown 3.999 Change <±3%/year Display overload „OL“ Sampling time 4 times/ sec Working temperature 0°C ~ 50°C working hummidity <80% relative hummidity Dimensions max 132mm x 60mm x 38 mm weight (approx) 150g Batteries 9V Battery time around 100 working hours |  | kom | 1 |  | 0 | 0 | 0 | 20%
9 | CST Apparatus | Type 304B CST, Weight 1.75 kg Dimensions max 33 x 26 x 5.5 cm |  | kom | 1 |  | 0 | 0 | 0 | 20%
10 | Transferpette | Transferpette 10-100µl |  | kom | 1 |  | 0 | 0 | 0 | 20%
11 | Transferpette | Transferpette 100-1000µl |  | kom | 1 |  | 0 | 0 | 0 | 20%
12 | Transferpette | Transferpette 500-5000µl |  | kom | 1 |  | 0 | 0 | 0 | 20%
13 | Transferpette | Transferpette 1000-10000µl |  | kom | 1 |  | 0 | 0 | 0 | 20%
14 | Fluidized bed reactor | TST-2L, stirring power 90W, speed 0-1300rpm, diameter of (PTEE) stirring bar φ6, no reflux condenser, no cap of heating bath |  | kom | 1 |  | 0 | 0 | 0 | 20%
15 | Water pump | Single-stage Pump, Electric 180W, Flow 60L/min, Size max 400*280*420mm, Water storage capacity 15L |  | kom | 1 |  | 0 | 0 | 0 | 20%
16 | Sušnica | Širenje toplote: Gravitaciono Temperaturni opseg: 50 °C - 250 °C Automatski alarmni sistem za prekoračenje temperature radi zaštite uzoraka Jedna vrata / sa uglom otvaranja od 180° Odstupanje prostorne temperature na 150 °C: ± 4 ° C Odstupanje temperature tokom vremena na 150 °C: ± 0,4 ° C Površina koju aparat zauzima: do 0,3 m2 Zapremina komore: 65 Lit Unutrašnje komore izrađene od nerđajućeg čelika otpornog na koroziju sa zaobljenim uglovima Dimenzije komore (Š x V x D): Maksimalno 328 x 480 x 415 mm Spoljašnja dimenzija (Š x V x D): Maksimalno 530 x 755 x 565 mm Broj polica: 2 Maks. nosivost police: 25 kg Napon / frekvencija: 230 V / 50 Hz Snaga / maks. napon: 1800 W / 7,9 A Potrošnja energije na 150 °C: 194 W |  | kom | 1 |  | 0 | 0 | 0 | 20%
17 | Autoklav | Autoklav zapremine 45mL, radna temperature 250°C, apsolutni maksimalni pritisak 1800psi, autoklav napravljen od čelika, sa zaštitom, teflonski uložak, PIN spanner wrench holding fixture. |  | kom | 2 |  | 0 | 0 | 0 | 20%
18 | Blok termostat | Digitalni termoblok sa dva bloka, Područje regulacije temperature: 5°C iznad ambijentalne do + 130°C, Područje podešavanja temperature: od + 15°C do + 130°C, Temperaturna stabilnost: ± 0,1°C, Temperaturna ujednačenost bloka: ± 0,1°C, Tip ekrana LED, Obezbeđenje od pregrevanja pomoću Termalnog osigurača, Tajmer sa zvučnim alarmom: 1 min – 72 sata, Vreme zagrevanja od + 25°C do + 100°C 15 minuta, Potrošnja električne energije: 300W, QB-50: 4 × 50ml, 50mm dubine, Ø29mm I QB-18: 12x Ø18mm, 50mm dubine. Poklopac. |  | kom | 1 |  | 0 | 0 | 0 | 20%
19 | Magnetna mešalica | Magnetna mešalica sa grejanjem, maksimalna zapremina mešanja 10 litara, Opseg brzine:100 – 1500 min-1,Opseg temperature: 50 – 500°C,Snaga:1000 W,Kontrola: LED, Materijal izrade ploče: staklo-keramika, Dimenzije ploče:200 x 200 mm, Klasa zaštite prema DIN EN 60529: IP 21, elektronski kontaktni termometar sa PID kontrolom, resest funkcijom, poseduje tri radna moda, temperaturni opseg: -50 do + 450°C, rezolucija:0.1K, tačnost: ± 0,2 K + tolerancija PT 1000 senzora DIN IEC 751 klasa A, devijacija: ± 0,5 K, težina: 0,2 kg, klasa zaštite IP 54 |  | kom | 1 |  | 0 | 0 | 0 | 20%
20 | Dispensete, analog III | 1-10 ml |  | kom | 1 |  | 0 | 0 | 0 | 20%
21 | Dispensete, analog III | 2.5-25 ml |  | kom | 1 |  | 0 | 0 | 0 | 20%
22 | Transferpette® S pipette, adjustable single channel | 100-1000 ul |  | kom | 1 |  | 0 | 0 | 0 | 20%
23 | Transferpette® S pipette, adjustable single channel | 500-5000 ul |  | kom | 1 |  | 0 | 0 | 0 | 20%
24 | Transferpette® S pipette, adjustable single channel | 1000-10000 ul |  | kom | 1 |  | 0 | 0 | 0 | 20%
25 | Bench top rack for 6 Transferpette® S single- pipettes |  |  | kom | 1 |  | 0 | 0 | 0 | 20%
26 | Sistem za membransku filtraciju sa vakuum pumpom | Model sa jednim radnim mestom, Levak od nerđajućeg čelika, membrane diametra od 47 mm, kapacitet levka 500 ml, u kompletu sa poklopcem, 1 mesto, vakuum crevo - 2 metra, vakuum bocu od 2000 ml, 1 kom; mala i velika cevčica, gumeni čep, filter stopper), Vacusart, stopper za sprečavanje ulaska tečnosti u vakuum pumpu; Vakuum pumpa: dijafragmalna pumpa visokog učinka, Aluminijumska glava pumpe, ventil od čelika, dijafragma PTFE, vakuum pumpa sa dve dijafragme,100% bezuljna pumpa,radni kapacitet 31 l/min, završni vakuum 100 mbar, konektori za creva ID 9mm, klasa zaštite motora IP20, snaga motora: 170 W |  | kom | 1 |  | 0 | 0 | 0 | 20%
27 | pH metar | pH tester za brzo merenje pH vrednosti u tečnostima. Poseduje kombinaciju pH imerzione i temperaturne sonde za brzu i efikasnu temperaturnu kompenzaciju. Kalibracija aparata se vrši u dve tačke, tj. pH = 4,00 i pH = 7,00. automatska temperaturna kompenzacija, opseg merenja od 0-14pH, preciznost ±0,02pH, rezolucija ±0,01pH, opseg temperature od 0-60C, preciznost ±0,4ºC, rezolucija 0,1ºC |  | kom | 1 |  | 0 | 0 | 0 | 20%
28 | Inkubator šejker | Inkubator šejker sa platformom za 6x1000ml erlenmajere, Područje podešavanja temperature: od + 25 ° C do + 80 ° C, Područje regulacije temperature: 10 ° C iznad ambijentalne do + 80 ° C, Raspon kontrole brzine: 50–250 o / min (priraštaj 10 o / min), Temperaturna stabilnost: ± 0,5 ° C, Rezolucija podešavanja temperature: 0,1 ° C, Tačnost temperature na +37°C: ± 0,5 ° C, Temperaturna ujednačenost na +37°C: ± 0,5 ° C, Orbita: 20 mm, Ekran: LCD, 2 × 16 znakova, Digitalno podešavanje vremena: 1 min – 96 sati / neprekidno (uvećanje od 1 min), Zvučni signal za tajmer: da, Maksimalno opterećenje: 8kg, Maksimalno neprekidno vreme rada: 30 dana, Dimenzije unutrašnje komore: maksimalno 460 × 400 × 310 mm, maksimalne dimenzije thermostata : Ukupne dimenzije maksimalno (Š × D × V): 590x525x510mm. Težina manja od 42 kg. |  | kom | 1 |  | 0 | 0 | 0 | 20%
29 | Mini laminarna komora | 1. Komora sa HEPA filterom, efikasnosti 99,995% ( H14;EN1822) 2. Vertikalna laminarna komora klase ISO 3, sa vertikalnim strujanjem vazduha 3 Za rad u PCR laboratorijama, za rad sa nepatogenim mikrobiološkim uzorcima, nukleinskim kiselinama i ćelijskim kulturama 4. Radna površina od nerđajućeg čelika AISI 304L, sa zaobljenim ivicama 5. Prednje staklo se lako podiže i omogućava nesmetano i lako čišćenje unutrašnjosti komore 6. Komora izrađena od čelika presvučena antimikrobskim premazom 7. Električna utičnica koja se nalazi u unutrašnjosti komore, DEHS test port 8. Na staklenim zidovima sa leve i desne strane nalaze se 4 otvora ( po 2 sa svake strane) za moguće instalacije priključaka za gas, vacuum… 9. UV sterilizaciona lampa 10. LCD display sa dugmetom za uključivanje i isključivanje, svetlo, utičnicu, UV lampom. 11. Maksimalne spoljašnje dimenzije komore ne veće od 830x655x925mm 12. Svetlo 600 Lux 13. Nivo buke 53 dB |  | kom | 1 |  | 0 | 0 | 0 | 20%
30 | Električni uređaj za hlađenje uzoraka No ICE electronic ice bath - akn5.1 By Techne ili odgovarajuće | 1. Uređaj za hlađenje uzoraka, pogodan za vijale različite veličine i različite vrste vijala; 2. Aluminijumska komora za hlađenje, maksimalan kapacitet cca. 60 1.5 mL vijala; 3. Keramičke kuglice kao medijum za hlađenje (umesto leda), dijametar 1.5-2.5 mm, mogu da se autoklaviraju, 1.7 kg kuglica; 4. Neograničen broj hlađenja; 5. Temperaturni opseg od 0.0 °C do 40.0 °C; 6. Preciznost temperature ±1°C; 7. Prosečno vreme hlađenja od 20 do 4 °C cca. 60 min i prosečno vreme hlađenja od 20 do 0°C cca. 3-4 h; 8. Spoljašnje dimenzije (LxWxH) = maksimalno 190 x 240 x 225 mm i dimenzije komore za hlađenje (LxWxH) = maksimalno 155x108x77 mm; 9. Težina maksimalno 4.5 kg; |  | kom | 1 |  | 0 | 0 | 0 | 20%
31 | Automatski brojač za ćelije LUNA-II automated Cell Counter, LUNA-II™ Automated Cell Counter ili odgovarajuća | LUNA-II automated Cell Counter, LUNA-II™ Automated Cell Counter. Samostalni uređaj sa integrisanom optikom i softverom. Određivanje ukupnog broja ćelija, kao i broja živih i mrtvih ćelija u jednom ciklusu rada. Sistem ima autofokus koji radi po principu tečnog sočiva. Opseg veličine ćelija koje je moguće detektovati: 3 - 60um. 5. Vreme brojanja manje od 10 sec sa manuelnim fokusom, manje od 15 sec sa autofokusom. Opseg određivanja vijabilnosti ćelija: 0-100%. Integrisana memorija za čuvanje rezultata do 1.000 brojanja. Touch screen displej. Uz uređaj neophodno isporučiti - USB sa minimum 4GB memorije - Paket kaseta za minimum 100 brojanja - Kasetu za višestruku upotrebu - Boju za pripremu uzoraka Tryptan Blue u koncentraciji od 0,4%, pakovanje 2 x 1ml. Logos Biosystems, L40002 |  | kom | 1 |  | 0 | 0 | 0 | 20%
32 | Vakum pumpa - Diaphragm vacuum pump ROTILABO CR-MV100 ili odgovarajuće | Suitable for vacuum filtration, vacuum drying and solid phase extraction. Delivery rate 10 l/min End vacuum (abs.) 100 mbar Meas. range 1000 - 1 mbar Graduation 50 mbar Overpressure 2,5 bar, Carl Roth, ECK4.1 |  | kom | 1 |  | 0 | 0 | 0 | 20%
33 | Automatski kontroler za serološke pipete ROTILABO ili odgovarajuće | Battery-powered pipette controller ROTILABO®, Za seroloske pipete, sa drzacem, dve brzine pipetiranja, Pipette holder made of PP, autoclavable. Replaceable PTFE membrane filter prevents contaminations. Ergonomically shaped housing for fatigue-free working UV-resistant housing material 2 pipetting speeds (High or Low Speed) Gravity mode, for non-motorised fluid drainage 20 h continuous duty, can be used while charging is in progress LED-based low battery indicator User friendly thanks to concave shaped finger pads With labelling space on the side Five coloured pipette holders for distinguishing between apparatus Additional pipette adapter for low pipetting volumes Weight approx. 200 g, Carl Roth, TC16.1 |  | kom | 1 |  | 0 | 0 | 0 | 20%
34 | Laboratorijski šejker IKA Hula Dancer ili odgovarajuće | Test tube shakers Hula Dancer series Hula Dancer digital, sa promenljivom brzinom od 1200 do 2800 rpm. Type Digital Movement type Circular, vibrating Shaking amplitude (stroke) 4,5 mm Speed display speed-adjustable in 9 increments W x D x H max 123 x 123 x 42 mm Weight max 0,7 kg Mains connection 100-240 V, 50/60 Hz, Carl Roth, KHC8.1 |  | kom | 1 |  | 0 | 0 | 0 | 20%
35 | Opema za vertikalnu elektroforezu i western blot MINI-PROTEAN ili odgovarajuće | Vertikalna elektroforeza za 1 do 4 gela sa modulom za vlažan transfer Broj gelova 1–4. Ćelija sa poklopcem sa kablovima za napajanje. Modul za puštanje gelova. Ukupni volumen pufera za 2 gela, 800 ml. Ukupni volumen pufera za 4 gela, 1.000 ml Tipično vrijeme rada za SDS-PAGE 35–45 min (pri konstantnoj frekvenciji 200 V). Gel format Mini. Dimenzije (WxLxH) maksimalno 12x16x18cm. 2 kasete za držanje gela. Jastučići. Sklop elektroda. Rashladna jedinica Bio Rad, 1658030 - PowerPac™ Basic Power Supply,Napajanje 10-300 V, potpuno podesivo u 1 V koracima. 4–400 mA, potpuno podesivo u koracima od 1 mA. 75 W (maksimalno). Konstantni napon ili konstantna struja s automatskim prelazom. 4 para izlaza. Tajmer od 1 min do 99 sati 59 min, u potpunosti podesivo. Ima pauzu / nastavak funkcije. 3-digit LED ekran. Otkrivanje preopterećenja / kratkog spoja. Otkrivanje curenja. No load otkrivanje. Zaštita od prevelikog napona. Zaštita od prekomerne temperature. Zaštita ulaza. Osigurač vruće i neutralno. Dimenzije (WxDxH), maksimalno cm 21 x 24,5 x 6,5 cm. Težina maksimalno 1.1kg, Bio Rad 1645050 - Mini-PROTEAN® Tetra Cell Casting Stand with Clamp Kit for Single Core System. Stalak za izlivanje gelova komplementaran sa Mini Protean vertikalnim elektroforetskim sistemom 1 postolje 2 okvira 2 gumice za sprečavanje curenja gela 5 špatula za oslobađanje gela Bio Rad, 1658052 |  | komplet | 1 |  | 0 | 0 | 0 | 20%
36 | Automatska pipeta BioPette ili odgovarajuće . | Automatska pipeta od 1 do 5 mL. Mehanicke, mogu da se autoklaviraju, promenljiva zapremina, nalezu Capp Expell nastavci od 5 mL Labnet P3960-5000A |  | kom | 1 |  | 0 | 0 | 0 | 20%
37 | Kulometrijski titrator | Opseg odredjivanja: 10 µg - 200 mg H2O Mogućnost sledećih titracionih modova: • Kulometrisjka KF titracija • Kulometrisjka KF titracija sa blank oduzimanjem • Blank određivanje • GLP: Validacija Kulometra • Rezolucija: 0.1 ug H2O • Preciznost: 10 ug...1000 ug H2O: ± 3 ug ; >1000 ug H2O: 0.3% min. Titrator za Kulometrijske Karl Fiescher titracije • Opseg odredjivanja: 10 µg - 200 mg H2O • Rezolucija: 0.1 ug H2O • Preciznost: 10 ug...1000 ug H2O: ± 3 ug ; >1000 ug H2O: 0.3% min. Mogućnost sledećih titracionih modova: • Kulometrisjka KF titracija • Kulometrisjka KF titracija sa blank oduzimanjem • Blank određivanje • GLP: Validacija Kulometra Karakteristike instrumenta: • Memorija instrumenta: Mogućnost pamćenja min.100 metoda • Statistika rezultata: srednja vrednost, relativna i apsolutna standardna devijacija (do 20 odredjivanja) • Automatsko kondicioniranje • Stop drift kao endpoint kriterijum • Odredjivanje drifta Komponente uredjaja: • Instrument sa LCD displej-om • Generatorska elektroda sa dijafragmom • Indikatorska elektroda • Odvojena tastatura za unošenje parametara metoda • Neophodan potrošni materijal za početak titracije Povezivanje sa drugim uređajima: • 2x RS232 konekcija • „Remote“ konekcija • Konekcija za magnetnu mešalicu • Mogućnost povezivanja na upravljački softver preko kompjutera • Mogućnost povezivanja na Karl Fischer peć Servisni zahtevi: • Ovlašćeni dobavljač dužan je da obezbedi kvalitetnu i efikasnu servisnu podršku • Neophodno je da serviser poseduje sertifikat o obuci o odrzavanju opreme kod proizvođača, kao I IQ/OQ sertifikat o obučenosti za izradu odgovarajućih protokola i dokumentacije • Ovlašćeni dobavljač dužan je da dostavi uputstvo za upotrebu na engleskom jeziku i servisne informacije uz instrument |  | kom | 1 |  | 0 | 0 | 0 | 20%
38 | Magnetna mešalica | Magnetna mešalica za Kulometar: Magnetna mešalica, kompatibilna sa Kulometrom Brzina mešanja: 200…1900 rpm Ambijentalna temperature: +5 - +45 °C Relativna vlažnost: ≤ 60% Materijal: Čelični emajlirani lim i PBT termoplastika |  | kom | 1 |  | 0 | 0 | 0 | 20%
39 | Kulometrijska posuda | Posuda za kulometrijsku titraciju sa odgovarajućim zaptivkama i cepovima • 80...250ml; Borsilikatno staklo 3.3. • Zaptivke na otvorima posude od PTFE folijeKablovi za generatorsku i indikatorsku kulometrijsku elektrodu • Držač za posudu Molekulsko sito: 250 g, veličina pora 0.3nm Septe: 5kom. 16mm, silikonska guma Zatvarač za posudu dužine 22mm, materijal Polipropilen Stakleni zatvarač za generatorsku elektrodu dužine 34mm Regensi za Kulometroijsku titraciju: • Reagens za Kulometrijsku titraciju 2x 500ml • Standardi za Kulometrijsku tiraciju 1 pakovanje |  | kom | 1 |  | 0 | 0 | 0 | 20%
40 | Automatska Pipeta 100-1000μl | 1 Kanalna Sa varijabilnom zapreminom Manualna Zapremina 100-1000μl Preciznost 0,002 |  | kom | 2 |  | 0 | 0 | 0 | 20%
41 | Automatska Pipeta 1000-5000 μl | 1 Kanalna Sa varijabilnom zapreminom Manualna Zapremina 1000-5000μl Preciznost 0,19% |  | kom | 2 |  | 0 | 0 | 0 | 20%
42 | Automatska Pipeta 20-200 μl | 1 Kanalna Sa varijabilnom zapreminom Manualna Zapremina 20-200μl Preciznost 0,23% |  | kom | 1 |  | 0 | 0 | 0 | 20%
43 | Mikro centrifuga | Rotator fiksni sa uglom Brzina 10000RPM Kompatibilne epruvete 8x0,2ml; 8x0,5ml; 8x1,5ml Maksimalni RCF 7500 x g Ugaoni rotator sa 8 pozicija Dimenzije maksimalno 175x156x125mm Težina maksimalno 1,4kg |  | kom | 1 |  | 0 | 0 | 0 | 20%
UKUPNA VREDNOST PONUDE BEZ PDV-A | [8 empty] | 0
IZNOS PDV-A | [8 empty] | 0
UKUPNA VREDNOST PONUDE SA PDV-OM | [8 empty] | 0
Rok isporuke, iznosi _________________ od dana zaključenja ugovora. ( rok isporuke ne može biti duži od 45 dana od dana zaključenja ugovora)
Način, rok (dinamika) i uslovi plaćanja: virmanski na račun ponuđača, sledećom dinamikom: - avans u iznosu 100% ugovorene vrednosti, uvećan za pripadajući porez, odmah po zaključenju ugovora, nakon što ponuđač dostavi menicu za povraćaj avansa, sa meničnim ovlašćenjem i pratećom dokumentacijom i avansni račun.
Rok važenja ponude, je_______ dana od dana otvaranja ponude (Rok važenja ponude ne može biti kraći od 90 dana od dana otvaranja ponuda)
Garantni rok: _______ meseci od dana izvršene zapisničke primopredaje (ne kraći od 24 meseca).
NAPOMENA: Za sva tražena dobra mogu se ponuditi i odgovarajuća. Dobra moraju biti u originalnom fabričkom pakovanju, upakovana u ambalaži i na način koji mora dobra da obezbedi od delimičnog ili potpunog oštećenja pri utovaru, transportu, pretovaru i uskladištenju. Mesto isporuke su laboratorije u objektima u sedištu Naručioca, Novi Sad, Trg Dositeja Obradovića 3. Iz objektivnih razloga Naručilac može da odredi i drugu lokaciju isporuke. Kontrola dobara vrši se prilikom primopredaje.
PONUĐAČ JE DUŽAN DA UZ PONUDU DOSTAVI ZA SVAKU STAVKU /PARTIJU, KATALOG ILI IZVOD IZ KATALOGA), SA NUMERISANIM BROJEM PARTIJE NA KOJU SE ODREĐENA SPECIFIKACIJA ODNOSI.
Potpis ovlašćenog lica ponuđača:
_____________________________________________________